## Strain Ext Cal
STRAIN EXTENSOMETER CALIBRATION
ASTM E-83, Class B-2 Calibration
Type 1 Extensometer System
[7 empty] | Ext. Asset No.: | 27206 |  |  |  | Calibration Reference Values
|  | Operator: | TRS |  |  |  | Translator: | 23845 |  | PreAmp: | 1805.95
|  | Test Station: | 60032 |  |  |  | Translator Resolution: | 5e-05 |  | PostAmp: | 1.5939
|  | Lab Temp (°): | 72 |  |  |  | Translator Cal. Due Date: | 2020-10-07 |  | Total Gain: | 2878.4506
|  | Humidity (%RH): | 35 |  |  |  | Gauge Pin: | 71004 |  | DeltaK: | 1.001
|  | Calibration Date: | 2019-11-04 |  |  |  | DVM: | 90325 |  | Excitation: | 6
|  | DUC Card: | S2-J4B |  |  |  | Unit System(in / mm): | inches |  | Excitation Freq.: | -
[7 empty] | Strain Mag: | 0.0025 |  | Phase (deg): | -
|  | Calibration Due Date: | 2020-02-02 |  |  |  | Gauge Length: | 0.5 |  | Zero Offset: | -
|  | Min. Strain Range (%): | 0.25 |  |  |  | Calibration Type: | New
[7 empty] | FS Measuring Range (+/-%e): | 5 |  | Calibration Approval:
Comments:
|  |  | Gauge R1 | Actual GL | Error (%) | Gauge R2 | Actual GL | Error (%)
|  | Gauge Verification: | 0.5 | 0.5 | 0 | 0.5 | 0.5 | 0 |  |  | 0
-10
-20
-30
-40
-50
-60
-70
-80
-90
-100
|  |  |  | Actual Voltage |  | Extension Reading (Cal units) |  | Fixed Error (unit/unit) |  | Relative Error (%)
Actual Extension (Calibration units) R1 | Actual Extension (Calibration units) R2 | Target Extension (Calibration units) | Target Strain (%) | Run 1 | Run 2 | Run 1 | Run 2 | Run 1 | Run 2 | Run 1 | Run 2 | Repeatability (%)
-0.025 | -0.025 | -0.026 | -5.2 | -10 | -9.9775 | -0.025 | -0.024944 | 0 | 0.0001125 | 0 | -0.21635 | 0.21635
-0.02 | -0.02 | -0.02 | -4 | -7.9916 | -7.9895 | -0.019979 | -0.019974 | 4.2e-05 | 5.25e-05 | -0.105 | -0.13125 | 0.02625
-0.015 | -0.015 | -0.015 | -3 | -6.0001 | -5.9962 | -0.015 | -0.014991 | -5e-07 | 1.9e-05 | 0.0016667 | -0.063333 | 0.065
-0.01 | -0.01 | -0.01 | -2 | -3.9933 | -3.9941 | -0.0099833 | -0.0099852 | 3.35e-05 | 2.95e-05 | -0.1675 | -0.1475 | 0.02
-0.005 | -0.005 | -0.005 | -1 | -2.0033 | -1.9977 | -0.0050082 | -0.0049943 | -1.65e-05 | 1.15e-05 | 0.165 | -0.115 | 0.28
-0.0025 | -0.0025 | -0.0025 | -0.5 | -1.0005 | -0.9986 | -0.0025013 | -0.0024965 | -2.5e-06 | 7e-06 | 0.05 | -0.14 | 0.19
-0.00125 | -0.00125 | -0.00125 | -0.25 | -0.5006 | -0.4994 | -0.0012515 | -0.0012485 | -3e-06 | 3e-06 | 0.12 | -0.12 | 0.24
-0.000625 | -0.000625 | -0.000625 | -0.125 | -0.249 | -0.2492 | -0.0006225 | -0.000623 | 5e-06 | 4e-06 | -0.4 | -0.32 | 0.08
0 | 0 | 0 | 0 | 0 | 0
0.000625 | 0.000625 | 0.000625 | 0.125 | 0.2509 | 0.2506 | 0.00062725 | 0.0006265 | 4.5e-06 | 3e-06 | 0.36 | 0.24 | 0.12
0.00125 | 0.00125 | 0.00125 | 0.25 | 0.5007 | 0.5013 | 0.0012518 | 0.0012533 | 3.5e-06 | 6.5e-06 | 0.14 | 0.26 | 0.12
0.0025 | 0.0025 | 0.0025 | 0.5 | 0.9998 | 1.0015 | 0.0024995 | 0.0025038 | -1e-06 | 7.5e-06 | -0.02 | 0.15 | 0.17
0.005 | 0.005 | 0.005 | 1 | 1.9978 | 2.0026 | 0.0049945 | 0.0050065 | -1.1e-05 | 1.3e-05 | -0.11 | 0.13 | 0.24
0.01 | 0.01 | 0.01 | 2 | 3.9935 | 4.0007 | 0.0099837 | 0.010002 | -3.25e-05 | 3.5e-06 | -0.1625 | 0.0175 | 0.18
0.015 | 0.015 | 0.015 | 3 | 5.9997 | 6.0068 | 0.014999 | 0.015017 | -1.5e-06 | 3.4e-05 | -0.005 | 0.11333 | 0.11833
0.02 | 0.02 | 0.02 | 4 | 8.007 | 8.0101 | 0.020018 | 0.020025 | 3.5e-05 | 5.05e-05 | 0.0875 | 0.12625 | 0.03875
0.025 | 0.025 | 0.025 | 5 | 10.012 | 10.013 | 0.02503 | 0.025032 | 6e-05 | 6.5e-05 | 0.12 | 0.13 | 0.01
Red highlighted cells are out of spec. data
Measurement Uncertainty Statement

## GPM
|  | STRAIN EXTENSOMETER CALIBRATION | [9 empty] | Strain Cal.
Click here to save the document. Don't forget to upload it on GPM right after.
Id Extensometer |  |  |  | Frame |  | Procédure |  |  | Date cal | Technician
10011 |  |  |  | 20026 |  | 500.1 & ASTM E-83
Model |  |  | Serial |  | Gauge Length |  |  | Unit System | Date due | Checker
632-53F-14 |  |  | 1432 |  | 12 |  | mm
Calibration Parameters
Translator |  | Translator Due Date |  |  |  | DUC Card |  | Probes Length |  | Full Scale Min/Max (%)
pgo |  | 2020-12-20 |  |  |  | S2-J2B |  | 85 |  | -5 |  | 5
PreAmp |  | PostAmp |  | Total Gain |  | DeltaK | Excitation |  | Fine Zero (V) |  | Shunt Cal (Ω)
850 |  | 2.6113 |  | 2219.58 |  | 1 | 7.5 |  | -2.123 |  | 450 M
Calibration Conditions
|  | Lab Temp. (°C) | [6 empty] | Lab Humidiy (%HR)
Validity
Cal Type |  | Cancel prev. Cal. |  |  | Compliant |  |  |  | Max (%) | Min (%) |  | Thermal Exp. Max
As Found |  | Yes |  |  | No |  |  | Range | 2 | -2 |  | 0.8
Comments: | [7 empty] | Sub-Range | 0.55 | 0.55 |  | 0
New prob TOP
Value
|  |  |  | Zero Position on Translator |  |  | Gauge Pin |  | Gauge Vérification (%)
|  |  |  | 400 |  |  | #5 |  | 0.01 | 0.25
Strain Target |  | Strain Step (%) |  | Screw Translator Position |  | Run 1 | Run 2 |  | Error 1 | Error 2 |  | Repeatability
(%) |  | (%) |  |  |  | (%) | (%) |  | (%) | (%) |  | (%)
0.11
0.11
0.11
0.11
0.11
0.11
3.2067 | 0.11 | 3.175 |  | 700
2.5717 | 0.11 | 2.54 |  | 640
1.9367 | 0.11 | 1.905 |  | 580
1.3017 | 0.11 | 1.27 |  | 520
0.63917 | 0.11 | 0.635 |  | 460
0.53333 | 0.11 | 0.52917 |  | 450
0.4275 | 0.11 | 0.42333 |  | 440
0.32167 | 0.11 | 0.3175 |  | 430
0.21583 | 0.11 | 0.21167 |  | 420
0.11 | 0.11 | 0.10583 |  | 410
0 | 0 | 0 | 0 | 400 | 400
0.18333 | 0.18333 | 0.10583 | 0.10583 | 410 | 410
-0.56083 | -0.56083 | -0.635 | -0.635 | 340 | 340
-1.3017 | -1.3017 | -1.3758 |  | 270 | 270
-2.0425 |  | -2.1167 |  | 200
-2.7833 |  | #N/A |  | #N/A
Red highlighted cells are out of spec. data
Note on J25: Must be smaller than Range

## Vis micro
pgo
[5 empty] | 23817 |  |  |  | 23818 |  |  |  | pgo
0.000127 |  |  |  | Translator Resolution: | 0.000127 |  |  |  | 0.0001 |  |  |  | 0.000127
vis | Indication mm | Corrigé V | Corrigé (mm) |  | vis | Indication mm | Corrigé V | Corrigé (mm) | vis | Indication mm | Relevé (mm) | Corrigé (mm) | vis | Indication mm | Corrigé V | Corrigé (mm)
200 | 0.254 | 0 | 0 | [5 empty] | 0 | 0 |  | 0 | 200 | 0.254
210 | 0.2667 | 0 | 0 | [5 empty] | 50 | 0.05 |  | 0.05 | 210 | 0.2667
220 | 0.2794 | 0 | 0 | [5 empty] | 100 | 0.1 |  | 0.1 | 220 | 0.2794
230 | 0.2921 | 0 | 0 | [5 empty] | 150 | 0.15 |  | 0.15 | 230 | 0.2921
240 | 0.3048 | 0 | 0 | [5 empty] | 200 | 0.2 |  | 0.2 | 240 | 0.3048
250 | 0.3175 | 0 | 0 | [5 empty] | 250 | 0.25 |  | 0.25 | 250 | 0.3175
260 | 0.3302 | 0 | 0 | [5 empty] | 300 | 0.3 |  | 0.3 | 260 | 0.3302
270 | 0.3429 | 0 | 0 | [5 empty] | 350 | 0.35 |  | 0.35 | 270 | 0.3429
280 | 0.3556 | 0 | 0 | [5 empty] | 400 | 0.4 |  | 0.4 | 280 | 0.3556
290 | 0.3683 | 0 | 0 | [5 empty] | 450 | 0.45 |  | 0.45 | 290 | 0.3683
300 | 0.381 | 0 | 0 | [5 empty] | 500 | 0.5 |  | 0.5 | 300 | 0.381
310 | 0.3937 | 0 | 0 | [5 empty] | 550 | 0.55 |  | 0.55 | 310 | 0.3937
320 | 0.4064 | 0 | 0 | [5 empty] | 600 | 0.6 |  | 0.6 | 320 | 0.4064
330 | 0.4191 | 0 | 0 | [5 empty] | 650 | 0.65 |  | 0.65 | 330 | 0.4191
340 | 0.4318 | 0 | 0 | [5 empty] | 700 | 0.7 |  | 0.7 | 340 | 0.4318
350 | 0.4445 | 0 | 0 | [5 empty] | 750 | 0.75 |  | 0.75 | 350 | 0.4445
360 | 0.4572 | 0 | 0 | [5 empty] | 800 | 0.8 |  | 0.8 | 360 | 0.4572
370 | 0.4699 | 0 | 0 | [5 empty] | 850 | 0.85 |  | 0.85 | 370 | 0.4699
380 | 0.4826 | 0 | 0 | [5 empty] | 900 | 0.9 |  | 0.9 | 380 | 0.4826
390 | 0.4953 | 0 | 0 | [5 empty] | 950 | 0.95 |  | 0.95 | 390 | 0.4953
400 | 0.508 | 0 | 0 | [5 empty] | 1000 | 1 |  | 1 | 400 | 0.508
410 | 0.5207 | 0 | 0 | [5 empty] | 1050 | 1.05 |  | 1.05 | 410 | 0.5207
420 | 0.5334 | 0 | 0 | [5 empty] | 1100 | 1.1 |  | 1.1 | 420 | 0.5334
430 | 0.5461 | 0 | 0 | [5 empty] | 1150 | 1.15 |  | 1.15 | 430 | 0.5461
440 | 0.5588 | 0 | 0 | [5 empty] | 1200 | 1.2 |  | 1.2 | 440 | 0.5588
450 | 0.5715 | 0 | 0 | [5 empty] | 1250 | 1.25 |  | 1.25 | 450 | 0.5715
460 | 0.5842 | 0 | 0 | [5 empty] | 1300 | 1.3 |  | 1.3 | 460 | 0.5842
470 | 0.5969 | 0 | 0 |  | 0 | 0 | -0.113 | -0.00678 | 1350 | 1.35 |  | 1.35 | 470 | 0.5969
480 | 0.6096 | 0 | 0 |  | 100 | 0.127 | -6.3388 | -0.38033 | 1400 | 1.4 |  | 1.4 | 480 | 0.6096
490 | 0.6223 | 0 | 0 |  | 163.8 | 0.20803 | -4.98 | -0.2988 | 1450 | 1.45 |  | 1.45 | 490 | 0.6223
500 | 0.635 | 0 | 0 |  | 200 | 0.254 | -4.2122 | -0.25273 | 1500 | 1.5 |  | 1.5 | 500 | 0.635
510 | 0.6477 | 0 | 0 |  | 211 | 0.26797 | -3.98 | -0.2388 | 1550 | 1.55 |  | 1.55 | 510 | 0.6477
520 | 0.6604 | 0 | 0 |  | 258.3 | 0.32804 | -2.981 | -0.17886 | 1600 | 1.6 |  | 1.6 | 520 | 0.6604
530 | 0.6731 | 0 | 0 |  | 300 | 0.381 | -2.1013 | -0.12608 | 1650 | 1.65 |  | 1.65 | 530 | 0.6731
540 | 0.6858 | 0 | 0 |  | 305.5 | 0.38798 | -1.987 | -0.11922 | 1700 | 1.7 |  | 1.7 | 540 | 0.6858
550 | 0.6985 | 0 | 0 |  | 352.8 | 0.44806 | -0.993 | -0.05958 | 1750 | 1.75 |  | 1.75 | 550 | 0.6985
560 | 0.7112 | 0 | 0 |  | 400 | 0.508 | 0.00033333 | 2e-05 | 1800 | 1.8 |  | 1.8 | 560 | 0.7112
570 | 0.7239 | 0 | 0 |  | 447.2 | 0.56794 | 0.997 | 0.05982 | 1850 | 1.85 |  | 1.85 | 570 | 0.7239
580 | 0.7366 | 0 | 0 |  | 494.5 | 0.62801 | 1.996 | 0.11976 | 1900 | 1.9 |  | 1.9 | 580 | 0.7366
590 | 0.7493 | 0 | 0 |  | 500 | 0.635 | 2.1137 | 0.12682 | 1950 | 1.95 |  | 1.95 | 590 | 0.7493
600 | 0.762 | 0 | 0 |  | 541.7 | 0.68796 | 2.993 | 0.17958 | 2000 | 2 |  | 2 | 600 | 0.762
610 | 0.7747 | 0 | 0 |  | 589 | 0.74803 | 3.993 | 0.23958 | 2050 | 2.05 |  | 2.05 | 610 | 0.7747
620 | 0.7874 | 0 | 0 |  | 600 | 0.762 | 4.2237 | 0.25342 | 2100 | 2.1 |  | 2.1 | 620 | 0.7874
630 | 0.8001 | 0 | 0 |  | 636.2 | 0.80797 | 4.99 | 0.2994 | 2150 | 2.15 |  | 2.15 | 630 | 0.8001
640 | 0.8128 | 0 | 0 |  | 683.5 | 0.86804 | 5.986 | 0.35916 | 2200 | 2.2 |  | 2.2 | 640 | 0.8128
650 | 0.8255 | 0 | 0 |  | 700 | 0.889 | 6.332 | 0.37992 | 2250 | 2.25 |  | 2.25 | 650 | 0.8255
660 | 0.8382 | 0 | 0 |  | 730.7 | 0.92799 | 6.98 | 0.4188 | 2300 | 2.3 |  | 2.3 | 660 | 0.8382
670 | 0.8509 | 0 | 0 |  | 778 | 0.98806 | 7.973 | 0.47838 | 2350 | 2.35 |  | 2.35 | 670 | 0.8509
680 | 0.8636 | 0 | 0 |  | 800 | 1.016 | 8.4378 | 0.50627 | 2400 | 2.4 |  | 2.4 | 680 | 0.8636
690 | 0.8763 | 0 | 0 |  | 900 | 1.143 | 10.547 | 0.63282 | 2450 | 2.45 |  | 2.45 | 690 | 0.8763
700 | 0.889 | 0 | 0 |  | 1000 | 1.27 | 12.6553 | 0.75932 | 2500 | 2.5 |  | 2.5 | 700 | 0.889

## Form History
Date | Revision History | Author
2018-09-15 | new corporate form created | wtg
2018-12-04 | Final update
| added gauge zero verification | wtg
2019-01-02 | adj limit of error to 0.5% indicator
fix B40 and B44 contents
| opend for 0-10v calibration | wtg
2019-02-17 | reformat to adhere to E-83 presentation format | wtg
2019-05-03 | Rev 2.1 corrected formula error in repeatability
| added second actual extension column | wtg