## Strain Ext Cal
STRAIN EXTENSOMETER CALIBRATION
ASTM E-83, Class B-2 Calibration
Type 1 Extensometer System
[6 empty] | Ext. Asset No.: | 10011 |  |  |  | Calibration Reference Values | [6 empty] | # Extensometer | 10011
| Operator: | PGO |  |  |  | Translator: | 23817 |  | PreAmp: |  |  |  | Operator: | PGO |  |  |  | Translator: | 23817
| Test Station: | 20021 |  |  |  | Translator Resolution: | 0.000127 |  | PostAmp: |  |  |  | Test Station: | 20021
| Lab Temp (°): | 23 |  |  |  | Translator Cal. Due Date: | 2019-12-20 |  | Total Gain: | 2219.58 |  |  | Lab Temp (°): | 23 |  |  |  | Translator Cal. Due Date: | 2019-12-20
| Humidity (%RH): | 44 |  |  |  | Gauge Pin: | # 5 |  | DeltaK: | 1 |  |  | Humidity (%RH): | 44 |  |  |  | Gauge Pin: | # 5
| Calibration Date: | 2019-03-07 |  |  |  | DVM: | N/A |  | Excitation: | 7.5 |  |  | Calibration Date: | 2019-03-07
| DUC Card: | NA |  |  |  | Unit System(in / mm): | mm |  | Excitation Freq.: |  |  |  | DUC Card: | NA |  |  |  | Unit System(in / mm): | mm
[6 empty] | Strain Mag: | 0.06 |  | Phase (deg):
| Calibration Due Date: | 2019-06-05 |  |  |  | Gauge Length: | 12 | [10 empty] | Gauge Length: | 12
| Min. Strain Range (%): | 0.1524 |  |  |  | Calibration Type: | as found | [10 empty] | Calibration Type: | as found
[6 empty] | FS Measuring Range (+/-%e): | -5/5% |  | Calibration Approval: | [8 empty] | Full Scale Min/Max (%) | -5 | 5
Comments:
|  | Gauge R1 | Actual GL | Error (%) | Gauge R2 | Actual GL | Error (%)
| Gauge Verification: | 12 | 12.0012 | 0.01 | 12 | 12.03 | 0.25 |  |  | 0
[10 empty] | -10 |  |  |  |  | Calibration Reference Values
[10 empty] | -20 |  |  | PreAmp: |  |  |  |  | Test Range (%): | -1 | 1
[10 empty] | -30 |  |  | PostAmp: |  |  |  |  | Min. Cal +/- (%): | 0 | 1
[10 empty] | -40 |  |  | Total Gain: | 2219.58 |  |  |  | Thermal Expansion Max (%): | 0.8
[10 empty] | -50 |  |  | DeltaK: | 1
[10 empty] | -60 |  |  | Excitation: | 7.5
[10 empty] | -70 |  |  | Fine Zero (V):
[10 empty] | -80 |  |  | Shunt Cal (Ω):
[10 empty] | -90 |  |  | Excitation Freq.:
[10 empty] | -100 |  |  | Phase (deg):
[14 empty] | Run1 (%) | Run2 (%)
[13 empty] | Gauge Verification: | 0.01 | 0.25
|  |  | Actual Voltage |  | Extension Reading (Cal units) |  | Fixed Error (unit/unit) |  | Relative Error (%) |  |  | Zero Position on Translator |  | 400 |  |  | Optimal Gain: | 3755.1238 |  | New Gain : | 3755.1238 |  |  | Delta K: |  |  |  | FOR INFORMATION ONLY
Actual Extension (Calibration units) | Target Extension (Calibration units) | Target Strain (%) | Run 1 | Run 2 | Run 1 | Run 2 | Run 1 | Run 2 | Run 1 | Run 2 | Repeatability (%) | Strain Target | Strain Step (%) | Screw Translator Position | Run1 (%) | Run2 (%) | Run1 | Run2 |  | Run1 | Run2 |  |  | Run1 | Run2 |  |  | def visé | equivalent mm | dep avec décalage vis | valeur la plus proche mm | translation mm | def equivalent
[19 empty] | 2.0002 | [8 empty] | 1.6 | 0.192 | 551.2 | 589 | 0.24003 | 2.0002
[19 empty] | 2.0002 | [8 empty] | 1.4 | 0.168 | 532.3 | 541.7 | 0.17996 | 1.4997
[19 empty] | 2.0002 | [8 empty] | 1.2 | 0.144 | 513.4 | 541.7 | 0.17996 | 1.4997
[19 empty] | 2.0002 | [8 empty] | 1 | 0.12 | 494.5 | 494.5 | 0.12001 | 1.0001
0.24003 | 0.24003 | 2.0002 | 2.01 | 1.9902 | 0.1206 | 0.11941 | -0.0099525 | -0.010051 | -49.7563 | -50.25 | 100.0062 | 1.6 | 2.0002 | 589 | 1.005 | 0.99512 | 4417.6268 | 4461.4645 | 2.0002 | 1.7003 | 1.6836 | -0.29997 | -0.31668 |  |  | 0.10001 | -0.10001 | 0.8 | 0.096 | 475.6 | 494.5 | 0.12001 | 1.0001
0.17996 | 0.17996 | 1.4997 | 1.9822 | 1.9902 | 0.11893 | 0.11941 | -0.0050853 | -0.0050453 | -33.9099 | -33.6432 | 67.5532 | 1.2 | 1.4997 | 541.7 | 0.99112 | 0.99512 | 3358.4176 | 3344.9181 | 1.4997 | 1.6768 | 1.6836 | 0.17714 | 0.18391 |  |  | 0.074983 | -0.074983 | 0.6 | 0.072 | 456.7 | 494.5 | 0.12001 | 1.0001
0.12001 | 0.12001 | 1.0001 | 1.0011 | 1.0051 | 0.060064 | 0.060304 | -0.0049959 | -0.0049759 | -49.9529 | -49.7529 | 99.7059 | 0.8 | 1.0001 | 494.5 | 0.50053 | 0.50253 | 4434.9842 | 4417.3337 | 1.0001 | 0.84681 | 0.85019 | -0.15331 | -0.14993 |  |  | 0.050006 | -0.050006 | 0.4 | 0.048 | 437.8 | 447.2 | 0.059944 | 0.49953
0.059944 | 0.059944 | 0.49953 | 0.98107 | 0.97907 | 0.058864 | 0.058744 | -9e-05 | -0.0001 | -1.8017 | -2.0019 | 3.8035 | 0.4 | 0.49953 | 447.2 | 0.49053 | 0.48953 | 2260.3035 | 2264.9207 | 0.49953 | 0.82989 | 0.8282 | 0.33036 | 0.32867 |  |  | 0.024977 | -0.024977 | 0.2 | 0.024 | 418.9 | 447.2 | 0.059944 | 0.49953
0 | 0 | 0 | 0 | 0 | [7 empty] | 0 | 0 | 400 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 400 | 400 | 0 | 0
-0.12002 | -0.12002 | -1.0001 | -1.0171 | -0.98507 | -0.061024 | -0.059104 | 0.0049159 | 0.0050759 | -49.153 | -50.7528 | 99.9058 | -0.6 | -1.0001 | 305.5 | -0.50853 | -0.49253 | 4365.2152 | 4507.0197 | -1.0001 | -0.86035 | -0.83328 | 0.13978 | 0.16685 |  |  | -0.050006 | 0.050006 | -0.6 | -0.072 | 343.3 | 305.5 | -0.12002 | -1.0001
-0.17996 | -0.17996 | -1.4997 | [9 empty] | -1.2 | -1.4997 | 258.3 |  |  |  |  | -1.4997 | [6 empty] | -0.074983 | 0.074983 | -1.2 | -0.144 | 286.6 | 258.3 | -0.17996 | -1.4997
-0.24003 | -0.24003 | -2.0003 | [9 empty] | -1.8 | -2.0003 | 211 |  |  |  |  | -2.0003 | [6 empty] | -0.10001 | 0.10001 | -1.8 | -0.216 | 229.9 | 211 | -0.24003 | -2.0003
-2.0003
-2.0003
-2.0003
-2.0003
-2.0003
Red highlighted cells are out of spec. data
Measurement Uncertainty Statement

## Vis micro
23817
[5 empty] | 23817 |  |  |  | 23818
0.000127 |  |  |  | Translator Resolution: | 0.000127 |  |  |  | 0.0001
vis | Indication mm | Corrigé V | Corrigé (mm) |  | vis | Indication mm | Corrigé V | Corrigé (mm) | vis | Indication mm | Relevé (mm) | Corrigé (mm)
0 | 0 | 0 | 0 | [5 empty] | 0 | 0 |  | 0
0 | 0 | 0 | 0 | [5 empty] | 50 | 0.05 |  | 0.05
0 | 0 | 0 | 0 | [5 empty] | 100 | 0.1 |  | 0.1
0 | 0 | 0 | 0 | [5 empty] | 150 | 0.15 |  | 0.15
0 | 0 | 0 | 0 | [5 empty] | 200 | 0.2 |  | 0.2
0 | 0 | 0 | 0 | [5 empty] | 250 | 0.25 |  | 0.25
0 | 0 | 0 | 0 | [5 empty] | 300 | 0.3 |  | 0.3
0 | 0 | 0 | 0 | [5 empty] | 350 | 0.35 |  | 0.35
0 | 0 | 0 | 0 | [5 empty] | 400 | 0.4 |  | 0.4
0 | 0 | 0 | 0 | [5 empty] | 450 | 0.45 |  | 0.45
0 | 0 | 0 | 0 | [5 empty] | 500 | 0.5 |  | 0.5
0 | 0 | 0 | 0 | [5 empty] | 550 | 0.55 |  | 0.55
0 | 0 | 0 | 0 | [5 empty] | 600 | 0.6 |  | 0.6
0 | 0 | 0 | 0 | [5 empty] | 650 | 0.65 |  | 0.65
0 | 0 | 0 | 0 | [5 empty] | 700 | 0.7 |  | 0.7
0 | 0 | 0 | 0 | [5 empty] | 750 | 0.75 |  | 0.75
0 | 0 | 0 | 0 | [5 empty] | 800 | 0.8 |  | 0.8
0 | 0 | 0 | 0 | [5 empty] | 850 | 0.85 |  | 0.85
0 | 0 | 0 | 0 | [5 empty] | 900 | 0.9 |  | 0.9
0 | 0 | 0 | 0 | [5 empty] | 950 | 0.95 |  | 0.95
0 | 0 | 0 | 0 | [5 empty] | 1000 | 1 |  | 1
0 | 0 | 0 | 0 | [5 empty] | 1050 | 1.05 |  | 1.05
0 | 0 | 0 | 0 | [5 empty] | 1100 | 1.1 |  | 1.1
0 | 0 | 0 | 0 | [5 empty] | 1150 | 1.15 |  | 1.15
0 | 0 | 0 | 0 | [5 empty] | 1200 | 1.2 |  | 1.2
0 | 0 | 0 | 0 | [5 empty] | 1250 | 1.25 |  | 1.25
0 | 0 | 0 | 0 | [5 empty] | 1300 | 1.3 |  | 1.3
0 | 0 | -0.113 | -0.00678 |  | 0 | 0 | -0.113 | -0.00678 | 1350 | 1.35 |  | 1.35
100 | 0.127 | -6.3388 | -0.38033 |  | 100 | 0.127 | -6.3388 | -0.38033 | 1400 | 1.4 |  | 1.4
163.8 | 0.20803 | -4.98 | -0.2988 |  | 163.8 | 0.20803 | -4.98 | -0.2988 | 1450 | 1.45 |  | 1.45
200 | 0.254 | -4.2122 | -0.25273 |  | 200 | 0.254 | -4.2122 | -0.25273 | 1500 | 1.5 |  | 1.5
211 | 0.26797 | -3.98 | -0.2388 |  | 211 | 0.26797 | -3.98 | -0.2388 | 1550 | 1.55 |  | 1.55
258.3 | 0.32804 | -2.981 | -0.17886 |  | 258.3 | 0.32804 | -2.981 | -0.17886 | 1600 | 1.6 |  | 1.6
300 | 0.381 | -2.1013 | -0.12608 |  | 300 | 0.381 | -2.1013 | -0.12608 | 1650 | 1.65 |  | 1.65
305.5 | 0.38798 | -1.987 | -0.11922 |  | 305.5 | 0.38798 | -1.987 | -0.11922 | 1700 | 1.7 |  | 1.7
352.8 | 0.44806 | -0.993 | -0.05958 |  | 352.8 | 0.44806 | -0.993 | -0.05958 | 1750 | 1.75 |  | 1.75
400 | 0.508 | 0.00033333 | 2e-05 |  | 400 | 0.508 | 0.00033333 | 2e-05 | 1800 | 1.8 |  | 1.8
447.2 | 0.56794 | 0.997 | 0.05982 |  | 447.2 | 0.56794 | 0.997 | 0.05982 | 1850 | 1.85 |  | 1.85
494.5 | 0.62801 | 1.996 | 0.11976 |  | 494.5 | 0.62801 | 1.996 | 0.11976 | 1900 | 1.9 |  | 1.9
500 | 0.635 | 2.1137 | 0.12682 |  | 500 | 0.635 | 2.1137 | 0.12682 | 1950 | 1.95 |  | 1.95
541.7 | 0.68796 | 2.993 | 0.17958 |  | 541.7 | 0.68796 | 2.993 | 0.17958 | 2000 | 2 |  | 2
589 | 0.74803 | 3.993 | 0.23958 |  | 589 | 0.74803 | 3.993 | 0.23958 | 2050 | 2.05 |  | 2.05
600 | 0.762 | 4.2237 | 0.25342 |  | 600 | 0.762 | 4.2237 | 0.25342 | 2100 | 2.1 |  | 2.1
636.2 | 0.80797 | 4.99 | 0.2994 |  | 636.2 | 0.80797 | 4.99 | 0.2994 | 2150 | 2.15 |  | 2.15
683.5 | 0.86804 | 5.986 | 0.35916 |  | 683.5 | 0.86804 | 5.986 | 0.35916 | 2200 | 2.2 |  | 2.2
700 | 0.889 | 6.332 | 0.37992 |  | 700 | 0.889 | 6.332 | 0.37992 | 2250 | 2.25 |  | 2.25
730.7 | 0.92799 | 6.98 | 0.4188 |  | 730.7 | 0.92799 | 6.98 | 0.4188 | 2300 | 2.3 |  | 2.3
778 | 0.98806 | 7.973 | 0.47838 |  | 778 | 0.98806 | 7.973 | 0.47838 | 2350 | 2.35 |  | 2.35
800 | 1.016 | 8.4378 | 0.50627 |  | 800 | 1.016 | 8.4378 | 0.50627 | 2400 | 2.4 |  | 2.4
900 | 1.143 | 10.547 | 0.63282 |  | 900 | 1.143 | 10.547 | 0.63282 | 2450 | 2.45 |  | 2.45
1000 | 1.27 | 12.6553 | 0.75932 |  | 1000 | 1.27 | 12.6553 | 0.75932 | 2500 | 2.5 |  | 2.5

## Form History
Date | Revision History | Author
2018-09-15 | new corporate form created | wtg
2018-12-04 | Final update
| added gauge zero verification | wtg
2019-01-02 | adj limit of error to 0.5% indicator
fix B40 and B44 contents
| opend for 0-10v calibration | wtg
2019-02-17 | reformat to adhere to E-83 presentation format | wtg